## Slide 1
The Ultrasound Reception Process

## Slide 2
Learning Objectives
Blocked force as a "lumped" source on reception
Thevenin equivalent transducer/sources on reception
Measurement of receiving cabling properties
Characterization of receiver as impedance, gain
and measurement of these receiver parameters
Reception transfer function
Ultrasonic system model

## Slide 3
Transducer on reception
F
v
V
I
[TB]
Note: assumed directions
of v,I are now reversed from sound generation case so:
acoustic
sources ?
incident waves
scattered waves
receiving transducer
F , v

## Slide 4
Transducer on reception
pinc
pscatt
v (x,) n
F, force on
transducer
=
pinc
pscatt
FB , blocked
force
B
+
S
.
v n = 0
.
v (x,) n
.
original problem
problem I
problem II
pscatt
FS , force due to motion of
transducer face

## Slide 5
Fs
S
v (x,) n
.
problem II : transducer is acting as a transmitter
pscatt
FB()
Zr ()
a
F
v
[TB]
V
I
Transducer on reception

## Slide 6
FB = 2 Finc = twice the force due to the
incident waves only
B
v n =0
.
problem I
pscatt
2 Finc()
Zr ()
a
F
v
[TB]
V
I
pinc
Note: many authors assume the incident and scattered
waves can be treated as plane waves incident on a fixed planar surface. In this case, as we showed previously:
Transducer on reception

## Slide 7
Transducer on reception
Thevenin equivalent
reduced transducer and acoustic sources model
F
v
V
I
[TB]
V
I

## Slide 8
reduced transducer and acoustic sources model
V
I
Transducer on reception
under open
circuit conditions
but, recall
so
open-circuit voltage

## Slide 9
reduced transducer and acoustic sources model
V
I
Transducer on reception
To get the impedance, remove the blocked force source and
measure :
F
v
V
I
[TB]
but this is the same configuration as during
sound generation so

## Slide 10
Transducer on reception
reduced transducer and acoustic sources model
V
I

## Slide 11
V1
I1
V2
I2
Receiving cable – transfer matrix model
[R]
V1
I1
V2
I2
R11 = R22 so changing directions on
currents does not change [R]
transmission line model

## Slide 12
Receiving cable - experimental values
Cable - 5 ft
Fixture - 2.5 ft
Receiver
In immersion setups,
cabling and fixtures
holding the transducer
need to be characterized
0
5
10
15
20
0
0.5
1
1.5
0
5
10
15
20
-200
-100
0
100
0
5
10
15
20
0
50
100
0
5
10
15
20
-150
-100
-50
0
5
10
15
20
0
0.01
0.02
0
5
10
15
20
-110
-100
-90
-80
0
5
10
15
20
0
0.5
1
0
5
10
15
20
-200
-100
0
100
Frequency (MHz)
Frequency (MHz)
Amplitude
Phase
(Deg)
(Deg)
(Deg)
(Deg)
R22
R21 (1/)
R12 ()
R11

## Slide 13
Receiver - equivalent circuit
V0
I0
Z0 ()
e
VR = K() V0
Z0 () … receiver impedance
e
K() … receiver gain
receiver
output voltage

## Slide 14
Z0 () = V0 () / I0 ()
K() = VR() / V0()
Receiver - experimental
transducer
cable
amplifier
VR
I0
V0
current probe
e

## Slide 15
Ultrasonic System Components
current probe

## Slide 16
Receiver - experimental
0
0.5
1
1.5
2
2.5
3
3.5
4
8
9
10
11
12
13
14
15
Amplitude
Frequency (MHz)
-40
-20
0
20
40
60
80
100
120
140
Phase (Deg)
amplitude
phase
Receiver gain
gain: 20 dB
attenuation: 0 dB
measured with
a 2.25 MHz
transducer
Receiver settings:

## Slide 17
Receiver - experimental
Receiver impedance
0
0.5
1
1.5
2
2.5
3
3.5
4
0
200
400
600
800
1000
1200
-150
-100
-50
0
50
100
Frequency (MHz)
Amplitude ()
Phase (Deg)
amplitude
phase

## Slide 18
Reception process model
transducer +
sources
cabling
receiver

## Slide 19
Generation and Reception process models
Both the generation and reception process transfer functions
depend only on the transducer sensitivities and their electrical impedances:

## Slide 20
Ultrasonic System Model
acoustic/elastic
processes