## Slide 1
Transducer Sound Radiation

## Slide 2
Learning Objectives
Planar Immersion Transducer
on-axis near field - far field
radiation into solid-normal incidence, plane interface
diffraction correction
Spherically focused transducer
on axis field
focal spot size
radiation into solid-normal incidence, plane interface
diffraction correction

## Slide 3
Learning Objectives (continued)
Contact P-wave transducer on a solid
wave types present
directivity functions
Angle beam shear wave transducer
Overview of beam theories

## Slide 4
Ultrasonic Beam Models
Plane piston transducer radiating into a fluid
z,z'
x,x'
y,y'
r
x = (x,y,z)
y = (x',y',z' = 0)
dS
vz
infinite baffle
Can show that each area element that is in
motion acts like a source of spherical waves:

## Slide 5
If we let vz(x',y',) = v0() (piston model)
Adding up all such sources over the face of the transducer gives
the Rayleigh-Sommerfeld Integral

## Slide 6
For on-axis response of a circular transducer of radius, a
a
r
z
x
y
It can be shown that
the area element can be
written as dS = r dr d
direct wave
edge waves

## Slide 7
direct wave
edge wave
edge wave
baffle
Direct and edge waves as seen for a pulsed transducer

## Slide 8
0
0.5
1
1.5
2
2.5
3
3.5
4
0
0.2
0.4
0.6
0.8
1
1.2
1.4
1.6
1.8
2
z/N
Near field distance N = a2/
Maxima: z = N/(2m+1) m = 0,1,2,…
Minima: z = N/2n n = 1,2,3,...
a = 6.35 mm
f= 5 MHz
c = 1500 m/s
|p|
cv0
Example: for a 5 MHz,
1/2 in. diameter transducer radiating into water
N = 5 in. (approx.)
on-axis pressure:

## Slide 9
Paraxial approximation : a/z <<1
z
a
plane wave
C1(a,,z)
diffraction correction
on-axis pressure:

## Slide 10
0
0.5
1
1.5
2
2.5
3
3.5
4
0
0.2
0.4
0.6
0.8
1
1.2
1.4
1.6
1.8
2
z/N
0
0.5
1
1.5
2
2.5
3
3.5
4
0
0.2
0.4
0.6
0.8
1
1.2
1.4
1.6
1.8
2
z/N
exact result
paraxial result
|p|
cv0
|p|
cv0
on-axis pressure:

## Slide 11
function p = on_axis(zN, A,c,F)
% exact on axis pressure from a piston source
%radiating into a fluid. A is radius in mm, c the
%wavespeed of the fluid in m/sec, F the frequency in MHz,
% zN is the distance in the fluid divided by the near field
%distance a^2/lamba (lamba is the wavelength)
al= 1000*A*F/c; % a/lamba
ka = 2*pi*al; % ka for the transducer
kz = ka*al*zN;
ke = 2*pi*(al^2).*sqrt(zN.^2 + (1/al)^2);
p = exp(i*kz) - exp(i*ke);

## Slide 12
function p = par_on_axis(zN, A,c,F)
% paraxial axis pressure from a piston source
%radiating into a fluid. A is radius in mm, c the
%wavespeed of the fluid in m/sec, F the frequency in MHz,
% zN is the distance in the fluid divided by the near field
%distance a^2/lamba
al= 1000*A*F/c; % a/lamba
ka = 2*pi*al; % ka for the transducer
kz = ka*al*zN;
ke = ka./(2*al.*zN);
p = exp(i*kz).*(1 - exp(i*ke));

## Slide 13
MAT> z = linspace(.2, 4,500);
MAT> p = on_axis(z,6.35,1500,5);
MAT> plot(z, abs(p))
MAT> xlabel('z/N')
MAT> p = par_on_axis(z, 6.35, 1500, 5);
MAT> plot(z, abs(p))
MAT> xlabel('z/N')

## Slide 14
On-axis response at normal incidence to a plane interface
(paraxial approximation)
z1
z2
fluid
solid
dp
transmission coefficient
(at normal incidence)
(velocity/velocity)
same diffraction correction
expression as for a single fluid
displacement in the
solid
x

## Slide 15
z1
z2
a
d1
d2
x
d
"virtual" point where the edge
wave would arrive on-axis in the solid if cp2 = cp1

so
which gives, in the paraxial approximation
p1
p1
p2

## Slide 16
S
x
y
r
R
y =(x,y,0)
e
The far-field is usually defined as z > 3N - also
called the "spherical wave region"
Far-field beam of a planar piston transducer

## Slide 17
Define the 2-D spatial Fourier
transform of where
as
Then the far field pressure can be written as
angular beam profile
spherical wave

## Slide 18
a
b
x
y


In spherical coordinates
Rectangular Piston Transducer
z

## Slide 19
Example far-field pattern of a rectangular transducer
-6
-4
-2
0
2
4
6
-10
-5
0
5
10
x-axis distance
y-axis distance

## Slide 20
Circular Piston Transducer
a
e

z = 3N
z = 6N
|p|
cv0
-150
-100
-50
0
50
100
150
0
0.2
0.4
0.6
0.8
1
1.2
1.4
radial distance, mm

## Slide 21
MAT> ang = linspace(-10, 10,500);
MAT> [p,r] =far_field(ang,6.35,1500,5,3);
MAT> plot(r,abs(p), '--')
MAT> hold on
MAT> [p,r] =far_field(ang,6.35,1500,5,6);
MAT> plot(r,abs(p), 'red')
MAT> xlabel('radial distance, mm')
function [p, rho] = far_field(ang,A, c, F, RN)
% far_field computes the normalized far field pressure
% for a circular piston (omitting the exp(ikR) phase term)
% A is the radius of the transducer in mm, c the wavespeed
%in m/sec, F the frequency in MHz, and RN is
%the normalized radial distance in near field units.
% rho is the transverse distance (normal to z) in mm
ka = 2*pi*(1000*A*F/c);
al= 1000*A*F/c;
x = ka*sin(ang*pi/180);
rho =RN*(A*al)*sin(ang*pi/180);
p = -i*(ka/(al*RN))*besselj(1,x)./(x+eps*(x ==0));

## Slide 22
Spherically Focused Piston Transducer Radiating
Into a Fluid
O’Neil Model
uniform
velocity, v0
Sf ... spherical surface
a
R0

## Slide 23
For x on the central axis
dS = r dr d/q0
q0 = 1 - z/R0
z
h
a
R0
x
re
r

## Slide 24
a = 6.35 mm
R0 = 76.2 mm
f= 10 MHz
c = 1480 m/s
Geometric focus
on-axis pressure versus z/R0:
True focus
0
0.5
1
1.5
2
2.5
3
3.5
4
0
2
4
6
8
10
12
z/R
|p|
cv0

## Slide 25
function p = focused_on_axis(zR, A,c,F,R)
% on axis pressure of a spherically focused probe
% as a function of the normalized distance, zR = z/R
%A, radius of the transducer in mm. R , focal length in mm.
%c, the wave speed in m/sec, and F the frequency in MHz
al=1000*A*F/c;
ka=2*pi*al;
zN=(R/A)*(1/al)*(zR);
kz=ka*al*zN;
kR=2000*pi*F*R/c;
kh=kR-sqrt(kR^2-ka^2);
kre=sqrt((kz-kh).^2+ka^2);
p = (exp(i*kz) -exp(i*kre))./(1-kz./kR);
MAT> z=linspace(.2,4,500);
MAT> p = focused_on_axis(z,6.35,1480,10,76.2);
MAT> plot(zr,abs(p))
MAT> xlabel('z/R')

## Slide 26
Paraxial Approximation
plane wave
diffraction correction
C1(a, z, R0 ,)
q0 = 1 - z/R0
on-axis pressure:

## Slide 27
|p|
cv0
|p|
cv0
exact
paraxial
on-axis pressure:
0
0.5
1
1.5
2
2.5
3
3.5
4
0
2
4
6
8
10
12
z/R
0
0.5
1
1.5
2
2.5
3
3.5
4
0
2
4
6
8
10
12
z/R

## Slide 28
function p = par_focused_on_axis(zR, A,c,F,R)
% on axis pressure of a spherically focused probe,paraxial approx.
% as a function of the normalized distance, zR = z/R
%A, radius of the transducer in mm. R , focal length in mm.
%c, the wave speed in m/sec, and F the frequency in MHz
al=1000*A*F/c;
ka=2*pi*al;
zN=(R/A)*(1/al)*(zR);
kz=ka*al*zN;
kR=2000*pi*F*R/c;
qo=1-kz./kR;
p = (1-exp(i*ka*(A/R)*qo./(2*zR)))./qo;
MAT> z=linspace(.2,4,500);
MAT> p = par_focused_on_axis(z,6.35,1480,10,76.2);
MAT> plot(zr, abs(p))
MAT> xlabel('z/R')

## Slide 29
Another way to model focusing (in the paraxial approximation)
O
suppose on a planar aperture we have a spherical wave propagating (generated by a lens, for example)
then on the aperture we have a phase given approximately
in the paraxial approximation by

## Slide 30
Thus, suppose we use a Rayleigh-Sommerfeld model for a planar transducer and place this phase (in the paraxial approximation) in the integral:
Using the paraxial approximation and evaluating this integral exactly for x on the transducer axis gives for a circular transducer of radius a:
Similarly, off-axis values will also represent those from a focused transducer

## Slide 31
a
R0
x
Wave field in the plane at the geometric
focus of a spherically focused transducer
p/pmax
-10
-8
-6
-4
-2
0
2
4
6
8
10
-0.2
0
0.2
0.4
0.6
0.8
1
1.2
Wf
 … wavelength
F = R0 / 2a … transducer
F number
kay/R0


## Slide 32
On-axis response at normal incidence to an interface
(paraxial approximation)
z1
z2
fluid
solid
dp
transmission coefficient
(velocity/velocity)
same diffraction correction
expression as for a fluid
displacement
x

## Slide 33
p0 … pressure
ST
n
x'
D
d1
d1
s
p
’
x"
For bulk waves
D = |x' - x"|
Contact P-wave Transducer Model
solid
stress-free
surface

## Slide 34
Directivity functions
Kp , Ks
Kp
Ks
0
10
20
30
40
50
60
70
80
90
0
0.1
0.2
0.3
0.4
0.5
0.6
0.7
0.8
0.9
1
angle, degrees

## Slide 35
function [kp,ks] = directivity(ang, cp, cs)
% computes the directivity functions for a p-wave contact
%transducer. ang is angle in degrees, cp, cs are p- and s-wave %speeds
k = cp/cs;
angr = ang*pi/180;
x = sin(angr);
c =cos(angr);
g=(x.^2 -k^2/2).^2 + x.^2.*sqrt(1 - x.^2).*sqrt(k^2 - x.^2);
kp = c.*(k^2).*(k.^2/2 -x.^2)./(2.*g);
ks = (k*x <1).*c.*(k^3).*x.*sqrt(1 - k^2.*x.^2)./(2.*g);
MAT> x = linspace(0,90,200);
MAT> [kp,ks] = directivity(x, 5900, 3200);
MAT> plot(x, kp)
MAT> hold on
MAT> plot(x, ks)
MAT> xlabel('angle, degrees')

## Slide 36
For ' small Kp = 1, Ks = 0
Dp
R
Ep
Es
H
Dp … Direct P-wave
Ep … Edge P-wave
Es … Edge S-wave
H … Head wave
R … Rayleigh wave
Full set of waves :
integral
contains direct
and edge P-waves

## Slide 37
Angle Beam Shear Wave Transducer Model
P
SV , SH
P
SV
P
SV
solid
solid
P
P
SV
“fluid”
solid
Can replace elastic
wedge by equivalent
"fluid" (neglect shear
waves)
(small for incident
P-wave beyond
critical angle)

## Slide 38
Ultrasonic Beam Models
Numerically Intense Models
EFIT - Langenberg
Finite Elements - Lord
Boundary Elements - Rizzo
Edge Elements - Schmerr, Lerch
Surface Integral Models
Generalized Point Source -Spies
Rayleigh- Sommerfeld + High Freq. Asymptotics
- Schmerr, Lhemery, others
Line Integral Models
Boundary Diffraction Wave - Schmerr, Lerch

## Slide 39
Ultrasonic Beam Models
Other Basis Function Models
Gauss-Hermite Models - Thompson, Gray, Newberry,
Minachi, Margetan
Multi- Gaussian Models
Minachi, Spies, Schmerr and Rudolph, Cerveny (Seismology)

## Slide 40
Ultrasonic Beam Models
A few references – mostly paraxial models
Lerch, T.P., Schmerr, L.W. and A. Sedov,” Ultrasonic beam models: an edge element
approach,” J. Acoust. Soc. Am., 104, 1256-1265, 1998.
Thompson, R. B. and E.F. Lopez,” The effects of focusing and refraction on Gaussian ultrasonic beams,” J. Nondestr. Eval., 4, 107-123, 1984.
Newberry, B.P. and R.B. Thompson,” A paraxial theory for the propagation of ultrasonic beams in anisotropic solids,” J. Acoust. Soc. Am., 85, 2290-2300, 1989.
Schmerr, L.W., Rudolph, M., and A. Sedov,” Modeling ultrasonic transducer wave fields
for general complex geometries and anisotropic materials,” Review of Progress in
Quantitative Nondestructive Evaluation, D. O. Thompson and D.E. Chimenti, Eds., Plenum Press, New York, 19A, 953-960, 2000.

## Slide 41
Schmerr, L. W., Fundamentals of Ultrasonic Nondestructive Evaluation, Plenum Press, New York, 1998.
Spies, M., and M. Kroning,” Ultrasonic inspection of inhomogeneous welds simulated by Gaussian beam superposition,” Review of Progress in Quantitative Nondestructive Evaluation, D. O. Thompson and D.E. Chimenti, Eds., Plenum Press, New York, 18A, 1107-1114, 1999.
Minachi, A., Margetan, F.J., and R.B. Thompson,” Reconstruction of a piston transducer beam using multi-Gaussian beams (MGB) and its applications,” Review of Progress in Quantitative Nondestructive Evaluation, D. O. Thompson and D.E. Chimenti, Eds., Plenum Press, New York, 17A, 907-914, 1989.
Gengembre,N. and A Lhemery," Calculation of wide band ultrasonic fields radiated by water-coupled transducers into heterogeneous media," Review of Progress in Quantitative Nondestructive Evaluation, D. O. Thompson and D.E. Chimenti, Eds., Plenum Press, New York, 18A, 1107-1131, 1999.
Ultrasonic Beam Models