## Slide 1
Transducer Characterization

## Slide 2
Learning Objectives
Pulse-echo experimental determination of impedance,
sensitivity of commercial transducers
Experimental determination of effective radii and
focal lengths of commercial transducers

## Slide 3
Transducer impedance, sensitivity
To fully characterize these generation and
reception transfer functions, we need to be
able to obtain the transducers input electrical
impedance and their sensitivities

## Slide 4
Transducer impedance
Measurement of the transducer input impedance, Zin :
Pulser
Transducer
Current probe
V
I
water
Short cable

## Slide 5
Transducer impedance
0
5
10
15
20
0
500
1000
1500
0
5
10
15
20
80
90
100
110
120
130
140
150
Frequency (MHz)
Phase (Deg)
Amplitude ()
amplitude
phase
Example impedance measurements of two 5 MHz transducers:
transducer 1

## Slide 6
Transducer impedance
0
5
10
15
20
0
500
1000
1500
0
5
10
15
20
80
90
100
110
120
130
140
150
Frequency (MHz)
Phase (Deg)
Amplitude ()
amplitude
phase
transducer 2

## Slide 7
0
1
2
3
4
5
6
7
8
9
10
0
2
4
6
8
frequency
0
1
2
3
4
5
6
7
8
9
10
89
89.5
90
90.5
91
Transducer impedance
Impedance of a capacitor, Z =1/(-iC)
magnitude
phase

## Slide 8
pulser/receiver
Transducer sensitivity
Pulse-echo measurement setup for determining sensitivity
solid

## Slide 9
Transducer sensitivity
solid
a
b
1. measure voltage and current when transducer is radiating into the fluid but before any reflected waves have arrived

## Slide 10
Transducer sensitivity
2. do FFT of the measured voltage and current and relate to the voltage and current at the transducer by compensating for the cabling
Note: then the impedance of the transducer is just

## Slide 11
Transducer sensitivity
3. measure voltage and current generated by the waves reflected from the surface of the block:
solid
a
b

## Slide 12
Transducer sensitivity
4. do FFT of the measured voltage and current and relate to the voltage and current at the transducer electrical port by compensating for the cabling

## Slide 13
Transducer sensitivity
5. From these measurements and a knowledge of the acoustic/elastic transfer function for this setup we can obtain the sensitivity of the transducer from:
[Note: in all these division processes in the frequency domain, a Wiener filter is used to desensitize the process to noise]

## Slide 14
Example sensitivity calculated for a
5MHz planar transducer:
Transducer sensitivity
0
0.05
0.1
0.15
0.2
0.25
0.3
0.35
Amplitude (V/N)
f (MHz)
0
5
10
15
20
-200
-150
-100
-50
0
50
100
Phase (deg)

## Slide 15
Transducer sensitivity
Cabling effects need to be accounted for in determining the sensitivity:
0
5
10
15
20
0
0.1
0.2
0.3
Amplitude (V/N)
0
5
10
15
20
-200
0
200
400
f (MHz)
Phase (deg)
Compensation
No compensation
f (MHz)

## Slide 16
Transducer Effective Parameters
Use of manufacturer specifications for parameters such as
transducer radius and focal length in transducer models
do not always lead to good agreement with experimental
measurements.
a
R0
Thus, we need to determine experimentally "effective" values
for these parameters.

## Slide 17
Transducer Effective Parameters
Effective radius for an unfocused (piston) immersion probe
sample
FFT
5 MHz
0
0.5
1
1.5
2
2.5
3
3.5
4
0
0.2
0.4
0.6
0.8
1
1.2
1.4
1.6
1.8
2
|p|
cv0
z/N
zmin
plot 5 MHz component
versus z
z
find zmin
determine
aeff

## Slide 18
Effective radius and focal length for a spherically focused probe
0
0.5
1
1.5
2
2.5
3
3.5
4
0
2
4
6
8
10
12
|p|
cv0
zmin
zmax
where x is the root of:
Transducer Effective Parameters

## Slide 19
Effective transducer parameters determined experimentally:
Probes | Manufacturers Specs focal length radius (cm) (cm) | Estimated Parameters focal length radius (cm) (cm) | Center Frequency (MHz)
A | 7.62 0.476 | 13.47 0.451 | 10
B | 7.62 0.635 | 20.74 0.556 | 5
C | 7.62 0.476 | 7.45 0.469 | 15
These values should be independent of frequency but in
practice they do vary somewhat with the frequency component
used in their determination.
Transducer Effective Parameters