## Slide 1
Acoustic/Elastic Transfer Function

## Slide 2
Learning Objectives
Modeling the acoustic/elastic wave processes as
a transfer function
Blocked force concept
Special cases where the transfer function can be obtained analytically
Effects of material attenuation

## Slide 3
Acoustic/Elastic Transfer Function
All the complex wave propagation and scattering processes occurring in an ultrasonic measurement system can be characterized in terms of an acoustic/elastic transfer function
transmitted
force
received
"blocked" force

## Slide 4
pscatt
The Blocked Force on Reception
On reception the receiving transducer face moves and scattered waves are also generated
pinc

## Slide 5
x
y
incident
wave
reflected
wave
rigid
boundary
The Blocked Force on Reception
If the incident waves are treated locally like plane waves and diffraction effects are ignored, we can consider the
interactions of the incident waves with the transducer face as those of plane wave with an infinite planar interface. Now, consider the case where the interface is held rigidly fixed:
The Blocked Force on Reception

## Slide 6
The Blocked Force on Reception
x
y
incident
wave
reflected
wave
rigid
boundary
In this case it can be shown that at the interface the blocked force is just double the force due to the incident wave only as if the transducer were absent, i.e
pinc
S
where

## Slide 7
The Blocked Force on Reception
from vx = 0 at x = 0
x
(reflected pressure = inc. pressure)
Equation of motion:
or
Thus, also at x = 0
proof for normally incident waves

## Slide 8
Acoustic/elastic processes model
For a number of calibration setups the acoustic/elastic
transfer function can be obtained explicitly:
front surface reflection
back surface reflection
reflection from flat-bottom hole
reflection from flat cylinder

## Slide 9
Acoustic Transfer Function
Example: the acoustic/elastic transfer function of two aligned
circular piston transducers of different radii can be found.

## Slide 10
Special Cases
a = b :
Acoustic Transfer Function
At the high frequencies found for most NDE transducers
the above integral can be evaluated approximately to yield:

## Slide 11
a >> b :
a << b :
D >> a, b :
Acoustic Transfer Function
Other special cases:

## Slide 12
Acoustic Transfer Function
Attenuation of ultrasound in the fluid can be added
to this model:
 … attenuation coefficient
For water
1/mm
f… frequency in MHz
D … distance traveled

## Slide 13
Example acoustic transfer function (including fluid
attenuation effects)
Acoustic Transfer Function
D = 444 mm
a = 3.175 mm
0
5
10
15
20
0
0.05
0.1
0.15
0.2
0.25
0.3
0.35
0.4
0
5
10
15
20
-90
-85
-80
-75
-70
-65
-60
-55
-50
-45
-40
Frequency (MHz)
Phase (Deg)
Amplitude
amplitude
phase

## Slide 14
Acoustic Transfer Function
Goldstein, A., Gandhi, D.R., and W.D. O’Brien,” Diffraction
effects in hydrophone measurements,” IEEE Trans. Ultrasonics,
Ferroelectrics, and Freq. Control, 45, 972-979, 1998.