Table of Contents

2. Sample Derivatization and Basics of Mass Spectra Interpretation	3

Purpose of Derivatization	3

Methoximation	3

Silylation	4

Interpreting GC-MS Chromatograms – The Basics	4

Interpreting Mass Spectra – Fragmentation Patterns	5

Example 1: Interpreting the Mass Spectrum of Pyruvate	6

Example 2:  Interpreting the Mass Spectrum of Alanine	7

Example 3: Interpreting the Mass Spectrum of 2-Hydoxybutyrate	8

3. Software Installation	9

Xcalibur	9

Automated Mass Spectral Deconvolution and Identification System (AMDIS)	9

MSSearch 2.0	10

MET-IDEA	10

LIB2NIST	10

Python	10

Adding Python to the Windows Path	11

4. Importing Mass Spectral Libraries	11

Library Procurement	12

Accessing and Manipulating Libraries	12

Converting a .MSP Library to NIST Format	12

Loading NIST Libraries into Xcalibur and MSSearch 2.0	13

Importing a .MSP Library Directly into MSSearch 2.0	13

Creating a .MSP Library using MSSearch 2.0 Librarian	14

Converting MassLab-formatted Libraries (*.IDB) to NIST and .MSP Formats	15

5. Browsing Chromatograms in Xcalibur	16

File Conversion	16

Basic Viewing	17

Searching the Library	18

Background Subtraction	19

Examining a Chromatogram Scan by Scan	21

Extracted Ion Chromatograms	21

Plotting Multiple Chromatograms	22

Panning	23

Amplification	23

6. Chromatogram Deconvolution in AMDIS	24

Optimizing AMDIS Settings	24

Identification Parameters	26

Instrument Parameters	26

Deconvolution Parameters	27

Library Parameters	27

Scan Sets	28

Peak Deconvolution	28

Updating Your MS Library with Components Identified by AMDIS	32

7. Using MET-IDEA 2.0 for Integration of GC-MS Peaks	32

Generating an Ion-Retention Time (IRT) List	33

Parameter Setup	34

Refinement	37

IRT List Refinement	37

Peak Integration	37

Common Integration Problems	38

Adjusting the IRT List and Optimizing MET-IDEA Parameters to Correct for Integration Problems	39

8. Running Python Scripts	41

Basics of Running a Script	41

Useful DOS Commands	42

Altering Python's Path	42

Other Resources	43

9. Using Python Scripts to Process MET-IDEA Results	43

Processing Data Using Python	43

10. Hierarchical Clustering of Data Using GenePattern	47

11. Summary - Overall Flow for Processing Metabolomics Data	49

# 1. Introduction

In this tutorial we present a general workflow to guide researchers with no prior experience in metabolomics through the supervised analysis and visualization of GC-MS-based data. We begin with a brief description of sample derivatization and the types of fragmentation patterns one expects to see in the mass spectra of tert-butyldimethylsilyl (t-BDMS, sometimes abbreviated as TBS) derivatized compounds. We then spend the bulk of the tutorial describing the specifics of extracting quantitative metabolite readings from raw trace data. Most of the programs used in this tutorial are freely-available and detailed instructions on their use can be retrieved from their respective websites. Where relevant, we highlight strengths and limitations of each program. It should be noted that several of these software programs are limited to use in the Microsoft Windows environment, so we have restricted our tutorial to this platform.

Researchers unacquainted with MS technology may wish to first familiarize themselves with current instrumentation in the field. This would be prudent because choice of instrument greatly affects the quality and reproducibility of metabolomics data. There are many useful tutorials on the internet which describe the basics of MS, including different ionization sources, mass analyzers, and mass detectors. One such tutorial can be found on the website for the American Society for Mass Spectrometry (ASMS) at:

http://www.asms.org/whatisms/index.html

Another useful tutorial can be found on the Shimadzu website at:

http://www.ssi.shimadzu.com/products/product.cfm?product=fundamentals_gcms1

For more detailed information and additional references, we also direct the reader to our recent review on the use of mass spectrometry for metabolomics analyses: Mishur, R.J. and Rea, S.L., Mass Spec. Rev. 2011 Apr 28. doi: 10.1002/mas.20338.

# 2. Sample Derivatization and Basics of Mass Spectra Interpretation

## Purpose of Derivatization

Metabolites are chemically derivatized prior to GC-MS. Derivatization affords many benefits including enhancement of volatility, GC resolvability, and MS ionization. Also, an appropriately chosen derivative can facilitate parent ion identification in MS spectra. In our studies, we routinely employ two sequential derivatization steps – a substituent protection step using methoxylamine, and a subsequent silylation step.

## Methoximation

Reaction with methoxylamine HCl converts carbonyl groups to methoximes (R=N-O-CH3) (Scheme 1). Note that it does not derivatize carboxylic acid functional groups. Methoximation prevents ketones from undergoing enolization and inhibits formation of multiple derivatives during the silylation step. Also, following methoximation,-keto acids are protected from decarboxylation and reducing sugars are prevented from cyclization. (Note: syn- and anti- isomers of methoximes may be resolved by GC giving rise to two peaks.)

Scheme 1. Ketone Methoximation

Silylation

To increase the volatility of metabolites for gas chromatography, hydroxyls, carboxyls, thiols, and primary and secondary amines are converted to t-BDMS derivatives. We employ a mixture of N-methyl-N-(tert-butyldimethylsilyl) trifluoroacetamide (MTBSTFA)/1% tert-butyldimethylchlorosilane (TBDMCS), both compounds react with labile hydrogens to form bulky, hydrophobic silyl derivatives (R-O-t-BDMS, R-CO-O-t-BDMS, R-S-t-BDMS, R-NH-t-BDMS, R2-N-t-BDMS). Derivatization usually proceeds to completion and t-BDMS derivatives are 2-4 orders more stable to hydrolysis than corresponding trimethylsilyl (TMS) derivatives.

## Interpreting GC-MS Chromatograms – The Basics

One advantage of GC-MS over other techniques used for the collection of metabolomics data is that, as libraries of fragmentation patterns of compounds grow, less and less is one required to manually interpret spectral patterns of unknown compounds to decipher their structure. To provide the reader with a basic understanding of MS-based compound identification, we give a brief introduction to GC-MS spectra interpretation and some of the artifacts that can arise due to sample derivatization with t-BDMS. We also present three example mass spectra and explain how compound identity can be reconstructed from the fragmentation profiles.

When using GC-MS, researchers will be confronted with two levels of data. The gas chromatogram appears as a series of peaks separated over time, with each peak corresponding to a specific column retention time (RT). Depending on the complexity of the sample and the resolution of the GC column, peaks represent one or multiple compounds that have eluted into the MS, ionized, and fragmented into daughter ions. The mass spectrometer analyzes ions according to their mass–to-charge ratio (m/z), and a signal proportional to the number of each ion is recorded by the mass detector. Peaks on the gas chromatogram therefore represent total ion counts as a function of time; hence, the gas chromatogram is often referred to as the total ion current (TIC) chromatogram. When peaks contain more than one compound, mathematical algorithms are required to deconvolute the peaks in the TIC chromatogram into constituent metabolites. Associated with each point in the GC trace is information regarding the type and number of ions detected by the MS. Programs such as the Automated Mass Spectral Deconvolution and Identification System (AMDIS) specialize in this process by identifying patterns of daughter ions that concordantly increase then decrease across a selected peak of the GC. Information on the extracted ions and their retention times can then be exported into a separate program, such as MET-IDEA, for integration and subsequent quantification. It is at this point that most of the ‘heavy work’ of GC-MS data analysis resides and this step often requires significant supervision.

Among metabolites, volatilization efficiency can differ significantly, therefore the highest intensity peaks in the GC trace do not always correspond to the most abundant compounds in samples. In the MS, while ionization efficiency is fairly uniform, a compound may follow multiple fragmentation pathways with differing relative efficiencies. This can lead to different relative ion abundancies. If an investigator is only interested in identifying differentially-regulated metabolites between samples, then inclusion of an internal standard will be sufficient to normalize spectra and control for variations in sample processing and instrumentation, before multivariate statistical analyses are performed. For truly quantitative purposes based on GC-MS data, however, calibration curves must be prepared using purified standards. Ideally, standards should be run on the same instrument, as variations in instruments and instrument settings can affect quantification.

## Interpreting Mass Spectra – Fragmentation Patterns

Metabolites must be ionized following elution from the GC column, prior to analysis by MS. There are several different ionization techniques which are employed for MS, the choice of which will affect the amount of post-ionization fragmentation, and thereby the appearance of the mass spectrum. Only in the most gentle (“soft” ionization) of these techniques is the molecular (M+/-) ion seen. When GC is coupled to MS for metabolomics, an electron impact (EI) ionization source is most often used. This is a relatively “hard” ionization technique, resulting in fragmentation of the (M+) parent ion. MS fragmentation patterns are dependent upon the method used to derivatize samples, the type of mass analyzer, as well as the ionization source. For this reason collections of mass spectra (libraries) are compiled using standardized fragmentation settings and derivatization reagents.

In the examples that follow, attention of the reader is drawn to the fact that not all daughter ions in a given spectra are of equal intensity. As described above, metabolites may fragment along multiple paths, albeit with differing efficiencies. Instrumentation type and settings will affect these relative efficiencies. The following spectra were obtained using an electron impact ionization source (70eV) with a quadrupole mass analyzer. Libraries constructed in this manner are usually similar enough across other EI instruments to allow identification of a metabolite, although relative daughter ion abundances may not perfectly match. The reader should also note that spectra are scaled according to ‘relative ion abundance’, with the most abundant ion set at 100%.

## Example 1: Interpreting the Mass Spectrum of Pyruvate

Pyruvate has a single carbonyl group which is protected via reaction with methoxylamine. The subsequent silylation reaction produces a t-BDMS derivative. Typically in mass spectra of t-BDMS-derivatized compounds the mass ion that is most abundant, and often of highest molecular weight, results from fragmentation of the t-BDMS derivative, leading to a characteristic ion at M-57 corresponding to loss of -C(CH3)3. Often the M-15 ion (loss of -CH3) can be detected as well, but it is much less abundant than the M-57 ion.

M-15

M-57

## Example 2:  Interpreting the Mass Spectrum of Alanine

In amino acids, both the carboxylic acid and the amine are derivatized by t-BDMS; however, it should be noted that the t-BDMS group is too bulky to derivatize the nitrogen a second time.

In addition to the typical M-15 and M-57 peaks, fragmentation of amino acids also gives rise to other characteristic ions:

M = 89 - 2 (2H) + 230 (2 t-BDMS) = 317

M-15 = 317 - 15 (CH3) = 302

M-57 = 317 - 57 (C(CH3)3) = 260

M-85 = 317 - 85 (C(CH3)3 + CO) = 232

M-159 = 317 - 159 (O=C-O-t-BDMS) = 158

Note: In the environment of the mass spectrometer, molecular rearrangements will occur, which lead to nontrivial fragmentation patterns. The M-85 peak below is a good example.

M-85

M-159

M-57

M-15

## Example 3: Interpreting the Mass Spectrum of 2-Hydoxybutyrate

M = 104 - 2 (2H) + 230 (2 t-BDMS) = 332

M-15 = 317

M-57 = 275

M-57-28 = 247

The fragment corresponding to M-57-28 is characteristic of 2-hydroxy acids and is due to loss of -C(CH3)3 and CO.

M-57-28

M-57

M-15

Note: Ions at m/z of 73, 75, and 147 are commonly seen in our spectra and arise from rearrangement of the tert-butyldimethylsilyl group itself.

m/z 73 = Si-(CH3)3

m/z 75 = HO-Si-(CH3)2

m/z 147 = (CH3)3Si-O-Si(CH3)2

Presence of the latter fragment is indicative of the presence of at least two silylated groups in the metabolite of interest.

# 3. Software Installation

In this section we provide a brief introduction to several programs that we routinely use in our analysis of GC-MS data. In an effort to make our method accessible to the largest possible audience we have endeavored to use freely available software for analysis whenever possible. In addition, we provide download and installation instructions for each of these programs. Certain functions are performed better by one program than others, and more details on the use each of these programs for analysis of metabolomics data will be discussed in the forthcoming sections. Xcalibur, AMDIS, and MSSearch 2.0 function in conjunction with a user-supplied library of mass spectra. We have provided our library as a resource for investigators (Supplementary Files). Instructions on how to load this library into each program are provided in the following section.

## Xcalibur

Xcalibur is proprietary software from ThermoFisher (previously Finnegan), and is the instrumentation software that operates the gas chromatograph and mass spectrometer on Thermo systems. This software is useful for general chromatogram browsing, for manipulating file formats, and for cross-confirmation of extracted ion chromatograms (EICs) with MET-IDEA. It can also be used for quantitation purposes. While not available for free download, institutions which operate Thermo instruments may have additional licenses of the software that permit installation on stand-alone computers for data analysis. It is highly recommended that this program be installed if your institution has access to the software. However, access to Xcalibur IS NOT ESSENTIAL for the processing of metabolomics data. If data was obtained on a Thermo instrument, the native file format will be .RAW, and this file format can be converted into several other formats by the Thermo software. If one does not have Xcalibur installed, they will need to request their MS facility to provide the data in .CDF format for use with AMDIS and MET-IDEA. Most non-Thermo instruments are also capable of exporting data in this format.

## Automated Mass Spectral Deconvolution and Identification System (AMDIS)

AMDIS is a useful program for mass spectral quantitation; however, in our experience, the peak integration functions associated with this program are restrictive. We limit our use of this program to two functions: (i) deconvolution of closely eluting peaks in GC-MS chromatograms, and (ii) the generation of a list of unique ion masses and their associated retention times (ion retention times, IRTs). When a library of MS spectra is made available to AMDIS, this program can also identify library matches in the sample spectra.

The AMDIS software can be downloaded from http://chemdata.nist.gov/mass-spc/amdis/downloads/. The current version is 2.69 and runs on Windows. During installation the software may ask for the instrument data format; if running a Thermo instrument, select Xcalibur.RAW. Note that AMDIS can read many file types in addition to the .RAW format.

A comprehensive manual is also available at the above website. New users are strongly encouraged to consult this text.

## MSSearch 2.0

The program MSSearch 2.0 provides a user-friendly interface with MS libraries. Mass spectra of individual metabolites extracted by AMDIS can be manually screened against a library of choice. Spectra of unknown metabolites can also be loaded into any library for future record using this program. These compounds must of course be assigned a unique identifier.

MSSearch 2.0 can be downloaded from the US National Institute of Standards and Technology (NIST) at http://chemdata.nist.gov/mass-spc/ms-search/downloads/. Both MSSearch 2.0 and AMDIS are also available to download as a single file from http://chemdata.nist.gov/mass-spc/ms-search/. Download the demo file (NISTDEMO_08.exe). This should install both programs as well as a demo version of the NIST library. The Demo version is fully functional.

MSSearch 2.0 and AMDIS can also be found on Xcalibur installation disks. After Xcalibur is successfully installed, re-run the installation program and choose NIST Setup. This will automatically install AMDIS and MSSearch 2.0.

## MET-IDEA

To circumvent some of the limitations of AMDIS, we utilize the program MET-IDEA for peak integration. MET-IDEA uses an ion retention time list to extract metabolite peak information across multiple spectra.

The MET-IDEA software can be downloaded from http://bioinfo.noble.org/download/ after registering. A comprehensive manual is included with the software.

## LIB2NIST

This program is used in the absence of Xcalibur to change Library file formats into one compatible with MSSearch 2.0. LIB2NIST can be downloaded from http://chemdata.nist.gov/mass-spc/ms-search/Library_conversion_tool.html. Once unzipped, the program is immediately executable.

## Python

Python is a general purpose scripting language useful for many scientific purposes (see http://www.scipy.org/ and http://matplotlib.sourceforge.net/). We provide several scripts that expedite data formatting, normalization and transformation. Python can be downloaded from http://www.python.org/.

### Adding Python to the Windows Path

Unless you plan on running Python from the Python installation directory, its location will need to be added to the path so that Windows knows where to find it when it is summoned from the command prompt. The following instructions are for appending the Python path in Windows 7.

- Under the Start menu, right-click on the “Computer” tab and select “Properties” from the drop-down menu.

- Select “Advanced System Settings” on the left.

- Under the “Advanced” tab, click the button that says “Environment Variables…”

- Under “System Variables”, highlight the line that starts with “Path…” and then click “Edit…”

- At the end of the line that says “Variable value”, type a semicolon (;) followed by the path for Python (i.e. C:\Python27)

- Select “OK” to close all of the open dialogue boxes

# 4. Importing Mass Spectral Libraries

Libraries are compilations of mass spectra which have been collected for known compounds, as well as known unknown compounds (i.e. investigator-generated compounds that are yet to be identified). Below is an entry from a library which contains the compound name and mass spectra for 4-aminobutyric acid derivatized with tert-butyldimethylsilane. Pairs of numbers represent m/z ratios and intensities (ion counts), respectively, of contributing ions.

Name: 4-AMINOBUTYRIC ACID TBS 2X

DB#: 1

Num Peaks: 50

41 27; 45 27; 57 23; 59 82; 68 50;

72 15; 73 630; 74 66; 75 219; 76 11;

85 11; 86 31; 87 27; 88 27; 98 19;

99 19; 100 39; 101 11; 109 109; 114 11;

115 23; 117 19; 119 15; 129 15; 131 15;

132 15; 133 101; 142 262; 143 94; 144 82;

145 19; 147 991; 148 227; 149 121; 150 11;

160 15; 174 11; 200 35; 201 15; 216 43;

217 7; 258 289; 259 62; 260 27; 274 999;

275 246; 276 105; 277 15; 316 39; 317 11;

This example is presented in .MSP format, which is a common library file format compatible with AMDIS. Other file formats recognized by AMDIS include MassLab format (a group of files with the extensions .IDB, .IDI, .PDB, .PDI), as well as .MSL, .CSL, .ISL and NIST. Xcalibur and MSSearch 2.0 require library files to be in NIST format.

## Library Procurement

A number of free MS libraries for TMS-derivatized compounds are available for download from the Max Planck Institute in Golm, Germany (http://gmd.mpimp-golm.mpg.de/download/). Commercial libraries are also available, such as the NIST library (http://www.sisweb.com/software/ms/nist.htm), and cost around $US 2000. We have compiled our own t-BDMS library of compounds found in the C. elegans exometabolome by paring down a free t-BDMS library provided by the Max Planck institute to compounds unique to our samples. Unidentified compounds detected in our spectra (known unknowns) were appended to the library and assigned unique identifier numbers. Several spectra of purified compounds representing metabolites of intermediary metabolism have also been added to our library. This C. elegans-specific library is available for download as a Supplementary file.

Important Note: Mass spectral libraries can be comprised of spectra resulting from underivatized or derivatized compounds, depending on the analytical method used. If the derivatization method you are using doesn't match what the library creators were using, the library will be of little use. The Golm libraries were created with compounds derivatized by MSTFA (TMS) or MTBSTFA (TBS). Also, the spectra curated in a library depend on the instrument used to collect the data. Spectra collected on a quadrupole will differ from those collected on an ion trap. Many of the same daughter ions will be present, but their relative ratios may differ and hence library search functions may not recognize a metabolite if both pattern and intensity parameters are exploited. In our experience, MS libraries collected using the same type of instrument, and when derivatized with the same moiety, are portable across labs.

## Accessing and Manipulating Libraries

### Converting a .MSP Library to NIST Format

Xcalibur and MSSearch 2.0 each require MS libraries to be in NIST format. The NIST format consists of a directory containing 17 library-related files. For this reason we have chosen to provide our library in .MSP format. Instructions for converting the .MSP library to NIST format using the program LIB2NIST are given below:

- Download the library file (TBS_realab_cel_v2.0.MSP) provided in the Supplementary Files.

- Run LIB2NIST by double-clicking the executable. (Note: in Vista and Windows 7 you will    need to be logged on as “Administrator".)

- Navigate the browser to your .MSP file and open.

- Change the Output directory to the MSSearch 2.0 directory. This is typically located at C:\Program Files (x86)\NISTMS\MSSEARCH. Under “Output Format” select “NIST MS Library”.

- Highlight the Input Library and select “Convert”. The library should now appear in the list of NIST MS User Libraries. If you receive an error message “Could not create output directory”, choose a different output directory and then copy the created folder into your MSSEARCH folder. Ignore the error messages “Could not create alias file” and “Could not create signature file”.

### Loading NIST Libraries into Xcalibur and MSSearch 2.0

Once the library has been converted to NIST format, it will need to be loaded into Xcalibur and MSSearch 2.0:

- Open Xcalibur and select the Library Manager tab under the Tools menu. If the NIST Setup in Section 3 was done properly the NISTDEMO library should appear under the Manage libraries tab. You want to replace this library with the supplied library. Click on the button that says “Add” and navigate to the directory of your library.

- Xcalibur calls upon MSSearch 2.0. for library search and display functions. Open MSSearch 2.0 and select Tools → Library Search Options → Libraries (the new library can now be selected).

### Importing a .MSP Library Directly into MSSearch 2.0

MSSearch 2.0 should be able to directly import libraries that are stored in .MSP file format. We have observed that this program sometimes has problems importing large libraries so a work around is to manually convert the library file into a type that MSSearch 2.0 can automatically recognize using the program LIB2NIST, as detailed above. Following are instructions for importing a .MSP format library directly into MSSearch 2.0:

- Navigate to the library .MSP file.

GO CUBS!!!

- Select the Librarian tab at the bottom of the screen. Import the list of library entries using the MS Search 2.0 import function.

- Select the list and add to an existing library or create a new library and add the entries.

### Creating a .MSP Library using MSSearch 2.0 Librarian

MSSearch 2.0 is also capable of exporting libraries in .MSP format. This could be useful if you need to create a .MSP library from a NIST formatted library, e.g., for distribution, or for use in AMDIS:

- Open MS Search 2.0 and select the Librarian tab

- Select the "Export from libraries" button

The ID Number Search dialog box will appear

- Select the library you wish to export from the drop down menu

- Using the ID range in the Library Statistics box, type in the metabolite identity, or range of metabolite identities, that you wish to include in your exported library

- Hit Search. This will populate the spec list with the contents of the library

- Highlight all of the spectra

- Select the "Export" button

- Select a location and file name. Save as a .MSP file.

### Converting MassLab-formatted Libraries (*.IDB) to NIST and .MSP Formats

Several of the free libraries made available by the Max Plank Golm project are provided in MassLab format. Here we provide example instructions for converting these files into formats recognizable by MSSearch 2.0 and Xcalibur.

Converting MassLab to NIST format using Xcalibur:

- Open Xcalibur and select Library Manager under the Tools tab

- Select the Convert libraries tab and input the following:

Source Library Details:

Type: MassLab (*.idb)

Library: Browse to the .idb file of the extracted library

Process entries: Ignore

- Input the details of the library format you wish to generate

Target Library Details:

Type: NIST

Library: C:\PROGRAM FILES\NISTMS\MSSEARCH\some_name

(This address is the location where the library will be saved and is 					typically the directory where MSSearch 2.0 is installed, but it may vary. 				Add a name to the end of the file path.)

- Select "Add the library to the NIST software for use with Xcalibur" to make the library visible to the internal search algorithm of Xcalibur.

Once the NIST format library has been loaded into MSSearch 2.0 it can be exported in .MSP format (see above). This will be necessary to use the library in AMDIS.

# 5. Browsing Chromatograms in Xcalibur

Xcalibur is used to browse chromatograms generated by Thermo Scientific intruments (.RAW format) and to convert .RAW files to .CDF files for use by other programs (e.g. MET-IDEA). It is also the software that runs the GC-MS on Thermo platforms. This section provides a basic walkthrough of some of the functionalities of Xcalibur. Note that this is proprietary software, and depending on your institution’s licenses, you may or may not have access to it; however, it is not essential to the processing of metabolomics data. More detailed instructions on the use of Xcalibur can be found in the Xcalibur help files associated with the program. Note – the screenshots in this and the remaining sections reflect the sample data files, when necessary. Some screenshots are, however, only representative and do not reflect the actual data.

## File Conversion

- Open Xcalibur. You may receive an error message: “Failed to update the system registry” due to Windows security settings on some systems. Click OK. Place the sample data files eat-2_A.cdf and eat-2_B.cdf (see Supplementary Material) in a separate folder. Use the Xcalibur file converter to convert them to .RAW format by selecting File Converter under the Tools menu on the main menu screen of Xcalibur. The screen below will appear:

A. Select source files (.CDF)

B. Select destination file type (.RAW)

C. Select destination file location (keep all of the files in the same place)

D. Select "Add Jobs"

E. Select "Convert"

F. Exit the File Converter dialogue box

## Basic Viewing

- From the main menu of Xcalibur select Qual Browser.

- Open the sample file eat-2_1.RAW. The following screen will appear:

The top window shows the TIC (total ion current) chromatogram. This represents the combined signal of all masses detected by the mass spectrometer as a function of retention time. The bottom window will show mass spectra once a time range has been selected.

- To view the mass spectrum for a peak, select the push pin in the upper right hand corner of the TIC chromatogram, and zoom the TIC by dragging a box around the area of interest.

There are several buttons in the menu header which are useful for scaling or zooming/unzooming the chromatogram. If they are not present in your install of Xcalibur, they can be added by going to the View menu and selecting Customize Toolbar. These buttons can then be found under the Display tab. From left to right in the figure below, they are: Drag with Cursor, Reset, Zoom In X, Zoom Out X, Display All, Zoom In Y, Zoom Out Y, Auto Range, and Normalize (0-100%).

- Once zoomed, select the push pin for the mass spectrum window (1 in figure below). Select a time range across your peak by left-clicking and dragging (2). The corresponding mass spectrum will appear in the bottom window (3). This is an average across the time range.

1

3

2

## Searching the Library

Spectra can be searched against a mass spectral library directly from Xcalibur by right-clicking in the mass spectrum window and then selecting Search from the Library menu.

In this case valine was reported as the best match. In general, probabilities > 80% are likely matches, but it is always important to carefully compare your spectrum against the match. Differences between spectra may be from background, or they may represent real chemical differences between your compound and the library match. The relative intensities of the masses comprising your spectrum should be similar to the library, but variations between mass spectrometers will lead to differences. This is especially true if the library was curated on a different type of mass spectrometer from the one your data is collected on (e.g. ion trap vs. quadrupole).

As an alternative to using the built in library search function of Xcalibur, right-clicking on the mass spectrum and selecting Export to Library Browser from the Library menu will bring up NIST MS Search 2.0. Upon executing a search, the program will prompt you to “Overwrite” or “Prepend” the Spec List contents. The Spec List is a list of spectra you are examining. Overwriting it replaces the previous search with the current search. Prepending the contents performs the new search, but also keeps old searches in the list.

## Background Subtraction

For peaks which are in an area of high background (e.g. near the solvent peak), or have close (but resolved) neighbors, it may be beneficial to subtract out the background before performing a library search.

To subtract the background from a peak in the GC trace, you will need to have the pushpin for the mass spectrum selected. Right-click in the mass spectrum window and select 1 Range from the Subtract Spectra menu, then select a time range on the TIC chromatogram by left-clicking and dragging. For our purposes, selecting a small range (0.025-0.05 minutes wide) directly in front of the peak of interest is generally sufficient. The screen-shot below shows the mass spectrum of a peak at 22.15 minutes before base line subtraction. Note that the black line (circled in red in the chromatogram window) denotes the range which is being subtracted.

The mass spectrum of the peak eluting at 22.15 minutes after background subtraction is shown below. In this case, subtracting the background improved the library match score from 87.5% to 92.6%.

The background subtraction can be cleared using the right-click menu. It is also possible to subtract two ranges, which can be useful when examining a peak which is closely flanked on either side by neighboring peaks.

## Examining a Chromatogram Scan by Scan

As compounds elute from the GC, they are ionized and fragmented in the mass spectrometer. Ions are selected for detection based on their mass to charge ratio (m/z). The process of selecting and detecting ions is called scanning. Each scan records the abundance of ions within a pre-selected m/z range. For example, with our instrument in full-scan mode, ions in the range of 50-700 m/z are detected and their abundances recorded. The process of scanning takes time and amounts to about 0.5 seconds for a full-scan on our instrument. Therefore a peak which elutes in about 4 seconds will have about 8 scans performed. Because the scans take time, there is a trade-off between the scan range and the signal detected for each m/z. By setting the mass spectrometer to detect a key ion which is characteristic of a compound of interest, one can increase the amount of time the instrument is detecting that ion and achieve a significant boost in sensitivity. This is the basis of Selective Ion Monitoring (SIM).

Instead of examining the mass spectra averaged across a range, you can select one time point (scan) and use the arrow keys to move sequentially through the scans. This is useful when there is a suspicion that multiple compounds are co-eluting in a single peak. If a peak is pure, the mass spectra will be more or less consistent in terms of ion identities across all scans of the peak. If more than one compound is present, a spectral mixture may be observed in part of the peak, assuming the compounds do not perfectly co-elute. The mixture will appear as a subgroup of masses which rise and fall together independently of another subgroup. An example can be seen in the peak eluting at 26.00 minutes. Select a point at the beginning of the peak and scan through to the tail.

## Extracted Ion Chromatograms

Instead of examining the TIC chromatogram, a chromatogram representing a single m/z can be plotted. This is known as an Extracted Ion Chromatogram (EIC), and is particularly useful when looking for the presence of a particular compound in a chromatogram if the expected fragmentation pattern of the compound is known. Plotting one or two fragments should locate the compound.

- Select the pushpin corresponding to the TIC chromatogram (top window).

- From the Display menu, select Ranges. Alternatively, you can right-click on the TIC chromatogram to access the chromatogram ranges window. The first entry in the Table should be the chromatogram you are currently viewing.

- Select the unchecked box on the first empty line.

- Change the "Plot type" to Base Peak

- Enter the desired m/z into the Range(s) field and click OK. In this field it is possible to enter individual masses, a mass range, or multiple ranges (separated by commas). In the example shown, a m/z of 158.1 has been chosen.

Zoom into the peak at 22.15 minutes. Since this peak contains a compound with a prominent fragment of m/z 158 an extracted ion peak appears as well.

Un-zooming will show all peaks where a fragment corresponding to m/z = 158 exists.

The extracted ion chromatogram can also be used to parse apart a peak comprised of co-eluting compounds. For example, the elution profiles of different co-eluting compounds at 25.15 minutes in the sample dataset can be seen by inspecting the EICs corresponding to m/z = 231 and m/z = 186.

## Plotting Multiple Chromatograms

The following procedure shows how multiple chromatograms can be plotted simultaneously in Xcalibur.

- Select Ranges under the Display Menu

- Select the first unchecked box (second line)

- Change the "Plot type" to TIC

- Where it says "Raw File", click on the button with three dots. Browse to the sample data file eat-2_2.RAW and click OK. The new TIC chromatogram will now be plotted underneath the first.

Note: Extracted Ion Chromatograms can be created for each TIC chromatogram loaded. In the "Ranges" window, first highlight the TIC chromatogram which the EIC will be based upon. Select a check box on the next open line, and follow the procedure for creating an EIC.

## Panning

Another useful feature in Xcalibur is panning. This can be done by zooming in to a portion of the TIC chromatogram and selecting Drag with Cursor from the Pan submenu of the Display menu. This will allow you to pan through the chromatogram, and is useful when comparing multiple chromatograms. A button to perform this function may also be added to the toolbar (see the subsection “Basic Viewing” in this section).

## Amplification

Ranges of the chromatogram may be magnified without zooming by selecting the magnification power in the drop down menu and then selecting the magnifying glass icon next to it. Left-click and drag on the desired range to magnify. To remove magnification, select the crossed out magnifying glass icon and left-click and drag over the range to de-magnify.

# 6. Chromatogram Deconvolution in AMDIS

The resolving power of gas chromatography is not unlimited and invariably some compounds in a complex mixture will co-elute. The program AMDIS is able to separate (deconvolute) co-eluting peaks computationally. The resulting peaks are referred to as components. The AMDIS algorithm extracts component peaks by identifying m/z clusters that appear and disappear with the same temporal profile within the original peak. A model component (peak) is generated for each cluster such that the combined sum of all components regenerates the original peak shape. After deconvolution, the extracted spectrum of a component can be compared against a mass spectral library to provide an identity to the peak. For further reading there is an overview on the AMDIS website (http://chemdata.nist.gov/mass-spc/amdis/) as well as a link to the technical paper by S. E. Stein.

The AMDIS algorithm is very sensitive and has a tendency to over-fit data leading to multiple extracted components for medium or large sized peaks, even if they are pure. Though parameters can be adjusted to minimize the occurrence of these events, they cannot be eliminated without the cost of missing small peaks. When a chromatogram is complex and comprised of many compounds AMDIS may therefore generate multiple peak entries that redundantly describe the same compound. The peak identification function of AMDIS is very useful, but we do not rely on AMDIS for peak quantitation. Instead, we use AMDIS only to generate a peak list from which we subsequently assemble an ion-retention time list that serves as the basis for quantitation using the program MET-IDEA. Ultimately, metabolite quantitation relies on identifying a representative m/z and then extracting its chromatogram from the original spectra over a defined retention time interval.

## Optimizing AMDIS Settings

There are many settings in AMDIS which may be adjusted by the user. Included among these are settings which affect the resolution and sensitivity of the algorithm, and a minimum library match factor which determines how “close” an extracted spectrum must be to a given library entry to be considered a match. Optimizing these parameters will affect the performance of the program.

Most menu items in AMDIS will be inactive until a file is loaded into the program. Load the file eat-2_1.cdf (included with the Supplementary Material) – the screen below will appear:

The top panel shows the TIC chromatogram and the bottom panel shows the mass spectrum of a selected scan.

Set the analysis settings by selecting Settings from the Analyze menu

There are six tabs under Analysis Settings, which will be discussed individually below:

1. Identif. parameters affect how AMDIS matches spectra against a chosen library.

2. Instr. parameter settings pertain to the type of instrument used to collect data.

3. Deconv. parameters pertain to peak deconvolution settings.

4. Libr. parameters specify the library file to use for peak identification.

5. QA/QC are quality control parameters.

6. Scan Sets: This parameter is for GC-MS runs in which the mass range was altered for different segments of the run. It may be absent in earlier versions of MET-IDEA.

### Identification Parameters

Minimum match factor: When comparing spectra against a library, potential matches are assigned a value representing the quality of the match. Matches > 80 are generally accepted as reliable. Setting the threshold lower to about 60-75 has several benefits:

- Low abundance peaks may not result in a high match score but may actually match a library entry

- Differences between the instrument the samples were run on and the instrument the library was produced on may lead to lower match scores.

- Common fragments between co-eluting peaks, such as 73, 75, and 147 (which are derived from the derivatization reagent), may not be extracted, and will thus affect the library match score.

Multiple identifications per compound: Selecting this box allows a library compound to be matched to multiple chromatographic peaks and is useful if you expect isomers or if a compound has multiple sites where t-BDMS derivatives could attach. We recommend selecting this box for metabolomics experiments.

Show standards: This is only used if performing an analysis with internal standards which have been set in the software. We do not employ this function.

Only reverse search: A reverse search calculates a match score using only those ions which are present in the library entry and ignores any which are not, and then compares this value against the minimum match factor. Score values calculated in the normal way would be negatively affected by any peaks present in the sample spectrum that are not in the library entry. We do not employ this function. (Earlier versions of AMDIS do not have this option)

Type of analysis: The "Simple" analysis matches extracted peak data to library entries using mass spectra only. Other types of analyses (e.g. retention indexed data) are possible if the library is retention indexed.

### Instrument Parameters

Low m/z: This pertains to the instrument scan range, in this case the lowest m/z value. Leaving at "Auto" will read this information automatically from the data file.

High m/z: This refers to the highest m/z value in the scan range. It can also be set to “Auto”.

Scan Direction: Instrument scan direction. Leave at default.

Instrument type: Type of mass analyzer. Our data has been collected on a quadrupole, change this field so it reads “Quadrupole”.

Use scan sets: Select if multiple scan ranges were used during the GC-MS run. This feature may not be available in earlier versions of AMDIS.

Threshold: This applies a noise filter to the data. We do not employ this function.

Data File Format: This refers to the instrument file format. For processing our data, change this setting to “Xcalibur Raw File”.

Set Default Instrument: Clicking on this button will allow the user to select a default data file format. The default type of mass analyzer may also be selected by clicking on the “Details>>” button in the dialogue box that pops up. This may later be changed by returning to the Settings dialogue.

### Deconvolution Parameters

(A word on AMDIS nomenclature:  Peaks refer to GC peaks. AMDIS parses each GC peak to determine whether it likely contains a single metabolite or, instead, is comprised of multiple metabolites. In the latter case, the peak is mathematically deconvoluted into component peaks, each characterized by a mass spectrum containing two or more extracted ions - see Figure 1B in main article.)

Component width: This represents the number of scans across the "average" peak. This doesn't need to be precise and the default value of 12 is okay. If you have larger peaks it can be increased.

Omit m/z: This option allows m/z values to be excluded as model peaks. It may prove useful to exclude m/z values which are present in the spectra of most compounds to prevent errors in integrating closely eluting peaks. We do not routinely utilize this parameter in our studies.

Adjacent peak subtraction: When extracting a component's spectrum, AMDIS may mistakenly include one or more ions from neighboring peaks. By selecting ‘adjacent peak subtraction’ the number of such mistaken ions that can be removed can be specified. Set at zero only if the chromatogram is very clean. A value of one is typically sufficient.

Resolution: Increasing the resolution permits AMDIS to extract peaks that are closer together.

Sensitivity: Increasing the sensitivity allows AMDIS to identify smaller, broader peaks.

Shape requirements: This parameter alters how similar an extracted ion's shape must be to a seed ion’s shape (called the model ion by AMDIS) before it is added to the component spectra. This can generally be left at medium.

Note: For metabolomics experiments, we have observed that the two parameters which have had the greatest effect on data analysis are Resolution and Sensitivity.

### Library Parameters

MS libraries/RI data: This allows the user to view and select the library used for mass spectrum matching. The entry “Target Compound Library” is the base library used for identifying compounds. The other libraries are used for more advanced analyses. Set this to the path of your .MSP library by clicking the “Select New” button. The Target Compounds Library field displays the current setting.

QA/QC Parameters

Solvent Tailing: This monitors the amount of peak tailing. Solvent or peak tailing often results from problems with the sample injection step. It is recommended to leave this setting at the default value.

Column Bleed: This monitors the amount of stationary phase degradation products leaching from the column. High column bleeding is a sign that the column may be damaged, and results in high baseline readings as well as a reduced signal-to-noise ratio. High column bleed may also result in less reliable library hits. It is recommended to leave this setting at the default value.

### Scan Sets

This parameter is not used with our runs and can be left at the default value. This tab is absent in earlier versions of AMDIS.

## Peak Deconvolution

After adjusting all of the parameters in Settings, click on the Save button. You will receive a dialogue box that says “Parameters of analysis had been changed. Reanalyze?”. Select “Yes” to run the analysis. If you do not receive this dialogue, you can run the analysis by clicking on the Run button in the top left corner of the screen. Once the processing is finished, the following screen will appear:

Targets are library matches. Components are deconvoluted peaks. The currently selected target or component will be highlighted in red.

The chromatogram can be zoomed by left-clicking in the TIC window and dragging a box, and/or converted to a log scale to improve visualization by right-clicking and selecting “Log Scale for Chromatogram”.

When zoomed, the chromatogram can be panned using the arrow buttons. By right-clicking in the TIC window you can also unzoom, rescale, and turn autoscale on/off.

The remaining panels give information on the extracted component and library matches.

A

C

B

D

Panel A represents the model (or seed) peak and other intense ions of the component. (Note that on this plot, the TIC (white curve) and EIC of the most abundant ion are both scaled to a maximum abundance of 100%. This means that the TIC trace will not equal the algebraic sum of the extracted ion traces.)

Panel B shows the scan spectrum of the original peak overlaid with the component spectra extracted by AMDIS.

Panel C compares the extracted component spectra with a library match.

Panel D gives information on the component and library match.

Things you can do:

- The component ion panel (A) will contain the model ion peak, and traces are automatically adjusted for any background and/or adjacent peak subtraction. Using the right-click menu overtop of the component ion panel (A), select Show Component. The extracted ions of a component also will then be plotted in the panel. This is a combination of the model peak and any background and/or adjacent peak subtractions.

- Using the right-click menu overtop of the main (TIC) chromatogram, select Show Component on Chromatogram. The component will then be plotted on the chromatogram.

- An extracted spectrum can also be appended to NIST MSSearch 2.0. Select a component, then right-click over top of the component spectrum and select Go to NIST MS Program under NIST Library.

- An Extracted Ion Chromatogram can be plotted, similar to Xcalibur. Go to Select m/z in the Options menu and type in the desired m/z to be plotted in the chromatogram or component panel. Also, clicking a m/z in the spectrum panel will plot it in the chromatogram panel. Clicking a m/z displayed on the right of the chromatogram will remove it from the display.

- Adjust the analysis parameters and see how they affect the results. Note: If you are zoomed in on the chromatogram, the analysis will only be run on the subset of the chromatogram. This can be useful when adjusting parameters as it will take a shorter amount of time to run. The goal of this exercise is to extract a component for every visible peak without generating a bunch of spurious components. The library parameters can be adjusted as well. Try to maximize the number of library matches without producing a lot of false positives. The only way to do this is by trial and error with spot checking. Once you are satisfied with the settings, un-zoom and re-run the analysis on the full chromatogram.

- AMDIS results can be exported as a report by selecting Generate Report under the File menu.

A. Select a location and file name.

B. Select "Append to report file" to add current results to a pre-existing file if desired.

C. Select "Report all hits" to report all library matches for each extracted component. It is useful to leave this unchecked and include only the first (i.e. best) hit to avoid a lengthy report file.

D. The report file can be opened in Microsoft Excel. See the AMDIS documentation for a complete description of the values returned.

- Multiple chromatograms can be analyzed simultaneously by running a batch job. Select Create and Run Job from the Batch Job submenu under the File menu. Add data files to analyze and select a file to save the results to. Click on the Run button to process the data.

- Two chromatograms can be displayed simultaneously in AMDIS. Select File→ Open In → Active Window.

7. Upon closing AMDIS it will prompt you regarding deletion of result files. These are intermediate files which AMDIS creates when it processes a chromatogram, and they contain the component spectra and library matches it found (.ELU and .FIN files).  Keep these files. As described in the next section, MET-IDEA relies on them. The standard text format for mass spectral data consists of a m/z followed by an intensity which has been scaled relative to the base peak; the base peak being the most abundant m/z in the spectrum.

Note that a result file is different from a report file. Report files give a summary of each peak found (RT, library match, match scores, integrated peak areas, etc.) in a tabulated format which can be opened in a spreadsheet program such as Microsoft Excel. This file can be generated by selecting “Generate Report” under the File tab.

Note also that when comparing similar samples, e.g. C. elegans exometabolome samples, you should expect to more-or-less see the same set of metabolites, or a subset of detectable metabolites, across samples, albeit at varying abundances. Therefore, after running through the AMDIS deconvolution algorithm on a few samples, and appending unknown components to your MS library (see next section), AMDIS may only identify (deconvolute) a few unidentified components in subsequent analyses. It is therefore possible to manually inspect your deconvoluted spectra in AMDIS, taking note of the retention times of components not already in the target library, and a representative ion to describe those peaks. This information can then be manually entered into the IRT list (see “Generating an Ion-Retention Time List” in Section 7).

## Updating Your MS Library with Components Identified by AMDIS

In addition to identifying library matches (targets) in the deconvoluted TIC chromatogram, AMDIS will also deconvolute components which are not in the MS library. It may be desirable to append the MS library with extracted mass spectra of these now ‘known unknown’ components.

- Select the component of interest by highlighting the upside-triangle above the component at the top of the screen.

- Under the Library menu, select Build One Library

- A dialogue box will appear; click on the button that says “Add ->: ##.### min filename”

- Highlight the new entry and click on Edit

- Assign a unique ID to the compound. A good strategy is to number new/unknown peaks sequentially. We use a six-digit ID number, with a unique first digit for each user who regularly updates the library.

- Sometimes you will observe retention time drift in your TIC traces (often due to miscalibration), and this can usually be fixed by applying a constant offset to all peaks in the dataset. Take note of where your internal standards are running. If they are not at their usual locations, you may want to adjust the RT of your unknown accordingly. When satisfied, click Save.

- Your .MSP file has now been updated, and re-running the analysis in AMDIS should identify your unknown peak as a library match. Note that in order for Xcalibur and MSSearch 2.0 to recognize your new compounds, you will have to convert your updated .MSP library to NIST format using LIB2NIST and then add its location to the Xcalibur library manager.

# 7. Using MET-IDEA 2.0 for Integration of GC-MS Peaks

MET-IDEA extracts peak area data from raw data by using an ion & corresponding retention time (IRT) list, where each IRT pair is chosen to uniquely describe a compound that one wishes to quantify. For each compound, MET-IDEA tracks to the specified retention time interval and then attempts to locate the specified ion (m/z). Within a pre-defined RT interval surrounding the specified RT value the ion count of the specified ion is integrated. A full description of the methodology can be found at http://pubs.acs.org/doi/abs/10.1021/ac0521596 (Broeckling, C. et al., Anal. Chem. 2006, 78(13), 4334-4341). All GC-MS data files must be converted to .CDF format prior to using MET-IDEA. If you are using a Thermo instrument, you can do this using the file conversion feature of Xcalibur, as detailed in section 5. Readers without access to the appropriate file conversion software will need to ask their MS facility to provide data in the .CDF file format.

## Generating an Ion-Retention Time (IRT) List

Before any data can be processed with MET-IDEA an Ion-Retention Time list must be generated. The list can be made either automatically or manually, or most likely via combination of the two. The final list looks like this:

Columns are (left to right): m/z, retention time (RT), six digit metabolite ID number, response to choice of whether to use metabolite for RT calibration, response to choice of whether to use metabolite for m/z calibration. The last two columns allow retention times and m/z accuracy to be calibrated during the analysis.

To produce an IRT list automatically, a data file is first processed in AMDIS as described in detail above. MET-IDEA will use the result file generated by AMDIS to create the IRT list. For every component in the result file, MET-IDEA will extract the retention time and the model ion used to establish that component.

Since the MET-IDEA IRT list is based on AMDIS’ results, it is best to optimize the AMDIS parameters before generating the list.

As specified in detail above, the method to refine the AMDIS parameters starts with examining a small region of the chromatogram, and testing different parameters. Your goals are to get a single component for every real peak, limit the numbers of false positive peaks (noise), and avoid the creation of false negatives (information loss). Compromises will have to be made. Generate an IRT list using the sample data file eat-2_1.cdf.

Make sure to save the AMDIS result files (.ELU and .FIN) when closing the program. Save this in the same folder as the .RAW and .CDF data files.

## Parameter Setup

To set up the MET-IDEA parameters, go to Parameter Setup under the Tools menu in MET-IDEA. There are three tabs in the Parameter Setup window. We describe only the most pertinent parameters which we routinely alter for our analyses. See section 3.1.2 of the MET-IDEA manual for a complete description of all of the parameters.

Note: If parameter changes are not accepted, then the user parameter file is read/write protected and its security settings need to be changed. Go to the MET-IDEA program folder (C:\Program Files (x86)\MET-IDEA) and find the file named UserParam.txt. Right-click on the file and select properties. Under the “Security” tab click “Edit” and then select “Users”. Allow permission to write to this file by checking “Allow” in the box that says “Write”. Click OK until you exit the dialogue.

The first tab in the Parameter Setup corresponds to the Chromatography parameters:

Average peak width (in minutes), Minimum peak width, and Maximum peak width must be determined by examining your chromatogram in Xcalibur or AMDIS. Note that several parameters are multiplication factors acting on the average peak width. The Peak start/stop slope parameter determines the threshold at which the algorithm determines the boundaries of a peak and can be adjusted later when examining the MET-IDEA results. For analyzing the sample data set, input 0.15 for Average peak width, 0.5 for Minimum peak width, 3 for Maximum peak width, and 1.5 for Peak start/stop slope.

The second tab corresponds to instrument parameters of the Mass Spec:

Under Mass spectrometers select the type of mass analyzer used to collect the mass spectral data which is to be analyzed. Mass accuracy regards the number of decimal points to which the analyzer collected its data; examine the data file in Xcalibur to determine this value. Mass range (+/-) refers to the mass accuracy drift of the instrument. There must be enough range allowed to accommodate the normal variation of the instrument, but not so much as to allow other ions to be picked up. The drift is best determined by examining masses across many runs on the instrument as well as across the range of m/z collected. For us, a mass range of +/- 0.2 is usually sufficient. If, when processing data with MET-IDEA, you are suddenly missing several peaks from the final result file, that run may be outside the mass range or retention time range. Parameters and constraints will therefore need to be expanded, or the IRT list adjusted manually. A good way to check to see if you have a wide enough mass range is to compare the extracted spectra in MET-IDEA to the EIC in Xcalibur. If they don’t match, you will probably need to adjust this parameter.

The third tab is the AMDIS parameters:

The Exclude ion list is a list of ions which will be ignored by MET-IDEA during the IRT list generation process. These are typically common background ions or ions which are common fragments (73, 75, 147). Note that if one of these ions is contained in the IRT list, MET-IDEA will then choose the next best candidate ion for integration. Setting a Lower mass limit of 150 will exclude these and other non-unique ions typical of the lower range of m/z. Ions per component is the number of IRT pairs per peak you wish to create. For our purposes it is left at 1.

After saving the parameter settings select Start from the MET-IDEA menu. The following window will appear:

A. Select the .ELU file generated by AMDIS (in this section you will generate a .ION file and will use this for subsequent analyses). Deselect the calibration settings for the moment.

B. Change the directory of the data folder to the location of the provided sample data files.

C. Click OK. The program will now create an IRT list. It may take a while.

D. If library matches are available they will be listed in the IRT list, otherwise an "unknown" will appear. The IRT list should be visually inspected for anomalies, for example, multiple entries for a single large peak. This is evident by a series of very similar retention times and/or ions. Inspection of the TIC chromatogram in AMDIS will also provide an indication of where to expect trouble. If you see a lot of anomalies, you may have to try to optimize settings in AMDIS and re-run. The IRT list will be saved in spreadsheet format, so it is also possible to manually remove redundant entries in the list using a spreadsheet editor such as Microsoft Excel (see also “IRT List Refinement” later in this section).

E. Click OK. You will be prompted to save the file, save it as a .ION file. The peak areas will then be extracted from the three sample data files and the results will be displayed. Take a quick look through the results to get a feel for the data and where problems may be occurring. Look for misshapen peaks, missing peaks, high background, integration across closely eluting peaks, etc. (see “Common Integration Problems” in this section”)

At this point is recommended to establish a file naming system so that files are sorted chronologically and indexed to notebook entries. For example, starting every filename with date_time(24hr) in the format of: YYMMDD_HHMM will automatically cause them to be listed in the order in which they were generated. In addition to raw data files, our lab has generated over 2,000 files related to processing, normalization, clustering, visualization, etc. in four years, so staying organized is critical. A master database file containing information such as filenames, sample contents, food source, culture conditions, total protein content, sample collection date, sample run date, and retention times of internal standards (in case retention time drift is observed) will prove useful when trying to locate data at a later date.

## Refinement

Refinement in MET-IDEA is the bottleneck in processing metabolomics data, as it is the most time consuming step. To produce high quality data, two things need to occur. First, the IRT list must be cleaned up from its raw, unrefined state. Second, the MET-IDEA parameters need to be adjusted in order to optimize the peak integration.

### IRT List Refinement

The IRT list needs to have one entry per extracted compound. AMDIS may have reported several components for one peak, and redundant entries will need to be removed. After looking across several data sets at the extracted ion chromatograms of component peaks you will most likely see integration errors (see below) resulting from poorly chosen ions. New ions will need to be chosen manually by examining the mass spectrum of the peak in question. A good ion is unique for the desired peak (i.e. not found in any close neighbors) and exhibits baseline background in the region of integration (i.e. not in the column bleed or solvent tail). For more information, see the section titled “Adjusting the IRT List and Optimizing MET-IDEA Parameters to Correct for Integration Problems” (below).

### Peak Integration

An important consideration in MET-IDEA when using an EIC as a surrogate measure of a peak’s area, is the relative abundance of the chosen ion in the peak's mass spectra. For example, consider the relative abundances of the ions in the following spectrum:

By putting several ions into a MET-IDEA IRT list and calculating the integrated areas it can be seen that the choice of ion affects quantitation.

From the above data, there is a 50 fold difference in the intensity of m/z 158.1 and m/z 302.2, even though they represent the same metabolite concentration. If possible, an abundant ion should be chosen to get the best sensitivity and peak integration.

The above example also underscores a point we belabored in the main text, GC-MS is great for relative metabolite quantitation across samples, as long as identical measurement and IRT integration values are used. In no way, however, can an extracted and integrated component ion, measured using the methodology described in this tutorial, be equated with absolute levels of a particular metabolite.

### Common Integration Problems

A number of common integration problems which may be encountered while using MET-IDEA are shown in the panels below:

A

B

C

D

E

A. Incomplete integration, usually small broad peaks. Try adjusting the peak shape parameter and/or minimum peak width.

B. High background AND ion is present in neighboring peak. Change ion to one with lower background and that is absent in neighboring peak.

C. Peak tailing on small peaks. Choose ion with low background. Alter peak shape parameter.

D. Noisy peaks, low abundance. Choose more abundant ion. Increase sample concentration, lower split ratio, or use selective ion monitoring.

E. A well-integrated peak.

### Adjusting the IRT List and Optimizing MET-IDEA Parameters to Correct for Integration Problems

- Open up MET-IDEA and start an analysis.

- When prompted for an IRT list, select the ‘eat-2_working.ION’ (provided in Supplementary Material) – this file must be in the same folder as the .CDF data files.

Notes: In the MET-IDEA open file dialog, change the file type from .ELU to .ION to see your IRT list. If your raw files are not in .CDF format, see “File Conversion” at the beginning of this section.

- Select "Calibrate Retention Time"; "Shift by constant". Deselect "Calibrate Mass Data".

The Calibrate Retention Time setting is to correct for instrument drift between runs. In our experience there is generally little day-to-day variation with regards to retention time. However, when drift does occur, calibration can be accomplished by selecting one or more peaks to anchor the analysis on. The retention times will be determined for these peaks, and a correction factor will be calculated based on the retention time in the IRT list provided. While shifting the data by a constant is generally sufficient, linear regression correction, employing multiple peaks, should ideally be used in order to prevent one poorly shaped peak throwing off the whole calibration with an inaccurate retention time. For this reason we advise use of at least three internal standards, to provide common peaks across all datasets. It is best to choose compounds with retention times that span the chromatographic region of interest and that are not expected to be present in the samples of interest. The type of sample be analyzed will also dictate internal standard identity. For our exometabolome studies described in the main text, we used phenylpyruvate, norvaline, and 3,4-dimethoxybenzoic acid.

Mass accuracy typically does not need to be calibrated unless there is a problem with the instrument. This can be determined by pre-inspection of the data in Xcalibur. The MS instrument should be tuned prior to analysis to make sure that it is recording accurate masses.

- Click the OK button. This will cause the "ViewIonList" window to appear. Scroll down the list of IRT pairs. The peak at 31.51 minutes (m/z = 239.02) corresponds to the internal standard 3,4-dimethoxybenzoate and will be selected as a RT (RT) and m/z (ION) marker. The other two internal standards will also be selected. The choice of these ions for RT (RT) and m/z markers is specified in the .ION file.

While in AMDIS (or Xcalibur), spot check a few ion masses to check for mass accuracy. If the masses in your data file differ significantly from the masses in the IRT list, then MET-IDEA will not pick up the ion trace.

- Once satisfied, click OK. The calibration results will be displayed. Examine the RT adjustments to see if they are reasonable. They should be fairly consistent. If one sample is way out, then a problem occurred with that sample and it should be closely examined before processing it.

- Click OK and save the calibration settings (.CAL file); the integration results will then be displayed. Tab through the results while examining the chromatogram. Look for any obvious integration problems (see the subsection “Common Integration Problems”, in this section, above).

- Attempt to correct any integration problems, either by manually adjusting the ions of the IRT list or the parameters of MET-IDEA. You may not be able to fix everything. In this case, integration regions will have to be defined manually for improperly integrated peaks. This is done by selecting a peak and entering values into the Start and End fields and then clicking on Apply. Note that if the Apply button is not pressed then MET-IDEA will not reintegrate over the new defined integral region. When finished, click “OK”. You will be prompted to save your progress, enter a filename. This will save the data as two files (a .POS and .OUT file).

# 8. Running Python Scripts

A script is a short program which is fed to an interpreter, in this case Python, which executes the commands contained within the script. Scripts containing more than approximately 150 lines of code essentially function as full-fledged programs. Scripts are simply text documents and can be written in any plain-text editor such as Notepad, Wordpad, or a more advanced text editor which developers typically use. Microsoft Word cannot be used as it adds formatting characters to the file and these will interfere with the interpreter. The text editors geared towards developers have settings which help monitor the formatting of your script. The scripts included with the Supplementary Information were created with Notepad++ (http://notepad-plus-plus.org/). IDLE comes packaged with Python but it is a poor text editor, although it does allow you to run scripts directly from the editor, which is convenient for debugging purposes.

## Basics of Running a Script

- Python will need the location of the script and any necessary input/output files in order to run. Python can be run from the command line in the home directory, but this requires specifying the full paths to all required scripts and data and can quickly become tedious. It is easier to navigate to the directory which contains the script and data files (note that the script files must be in the same directory as the data files to be processed), and then it will be possible to simply use the file names in the command line, as Python will automatically look for them in the current directory. In Vista, a command line window can be opened in the current directory (where the data is) by holding the shift key, right-clicking, and selecting "Open command window here". This function can be added to Windows XP by downloading Power Toys from Microsoft (http://windows.microsoft.com/en-US/windows/downloads/windows-xp).

Note that before using Python you will need to add its location to the Windows path. See Section 3 above “Software Installation” for instructions on how to do this.

- Download the files test_1.py, test_2.py, and test_2_data.txt (provided in Supplementary Material).

- Open a Command prompt in the directory containing your script files as detailed in step 1. Alternatively, you can access the Command prompt by selecting “Run” from the Windows Start menu, and typing “CMD”. Navigate to the directory containing your script files (see Useful DOS Commands, below). A script can now be run by using the following format at the command prompt:

python scriptname

Type “python test_1.py” to run the script test_1.py

- Some scripts require additional arguments such as a filename for input data, or other settings. The argument would then look like this:

python scriptname filename1 setting1

Note: The scripts contained with this tutorial contain comments embedded in the files themselves that provide descriptions of the purpose of each script, required input file formats, and outputted file formats. Scripts can be opened in any of the above-mentioned text editors.

After opening the script test_2.py to read the descriptive comments, try running it by typing “python test_2.py test_2_data.txt” followed by a “fun_factor”

For the fun_factor argument, enter an integer proportional to your love of C. elegans.

## Useful DOS Commands

Since you will be running Python scripts from the DOS (Command) prompt in Windows, it is useful to know a few DOS commands.

cd /User/folder1/folder2/ 	changes working directory to desired path

cd .. 				moves up a directory

cd / 				moves up to top level directory

cd				prints full path of working directory

dir				lists directories located in current directory

## Altering Python's Path

An alternative to copying all scripts to the same folder as your data files is to create a directory for the scripts and then add this directory to Python's path (not the same as the Windows system path described earlier). Python will now look for the script in the designated directory.

- Open a command prompt

- Run Python and type:

Import sys

sys.path.append("/your/path/to/scripts")

- Type sys.path to see all current paths

In some versions of Windows you may need to run this as administrator.

## Other Resources

Python tutorial: http://docs.python.org/tutorial/

Python coding guidelines: http://bayes.colorado.edu/PythonGuidelines.html

# 9. Using Python Scripts to Process MET-IDEA Results

Before visualization or statistical analysis can be performed, MET-IDEA results must be formatted and normalized. This can be done either manually in Excel or by using the included Python scripts.

The general flow of data processing will be as follows:

Format Peak Names

↓

Normalize Data to an Internal Standard

↓

Normalize Data to Total Protein (see main text M & M for methodology)

↓

Remove Known Artifacts

↓

Convert Compound ID Numbers to Names

↓

Apply a Cutoff Filter to Remove Absent Peaks (optional)

↓

Format Data for Further Analysis

## Processing Data Using Python

Before beginning, you will need the following files:

- The output file from MET-IDEA (peaks by row/samples by column)

- A .txt file of the total protein values for each sample (see below for file format details, an example file is included with the Supplementary Files)

- A .txt file of known artifacts (included with Supplementary Files, 100906_artifact_list.txt)

- A .txt file of a database which contains compound ID numbers vs. names (included with Supplementary Files, 2011_12_19_IDs_to_names.txt). As you update your library with new compounds, you will need to update this list as well.

- Process sample .CDF files using MET-IDEA as described above (if you have already done this, you can reopen your result file in MET-IDEA by going to the View menu and selecting Result). When finished, transpose the result file by clicking on the Transpose button. It is also possible to transpose the data manually in Excel. Peaks should be in rows, and samples by column (this is a tab-delimited file and the scripts have been written to read and write this format).

- Save a copy of the transposed file in the folder where the Python scripts are located.

- Run the script called 1_peak_name_reformat.py on the data file (MET-IDEA.OUT file) to reformat the peak names. Again, as mentioned above, if the scripts are opened in a text editor there are usage notes in each file on how to run that script. Most scripts require at least the declaration of an input file and perhaps one other argument. Compare the result file with the input file. The following figure should give you an indication of what you are looking for. Note the examples in this section are for illustrative purposes only. They indicate the changes that are occurring in the dataset following processing by Python; however, they do not correspond to the data files which you will be processing.

- Run the script called 2_normalize_to_IS.py on the reformatted data file. This will normalize the data to the internal standard by dividing each peak by its respective internal standard. The values are then multiplied by 10,000,000 to return data to the original scale. For our datasets we use the peak area of 3,4-dimethoxybenzoic acid (TBS 1X) (compound ID 100080) for normalization, and this is set as the default in the script. An alternative compound ID can be entered as the second argument when running the script. The script displays the result of normalizing the internal standard to itself: it should read 10000000.0 in all files. This is an arbitrary number which we have set for the intensity of the internal standard. The purpose of displaying the result is to indicate that the script ran successfully. If it does not display this, there is something amiss. You can also check to see if the normalization ran correctly by opening the output file generated by Python. In the row for the internal standard (in our case, 3,4-dimethoxybenzoic acid, ID 100080) you should see the number 10000000 in each column.

Note: Numbers in Python can be decimals, known as floating point, or integers (also, long integers and complex - see documentation). It is important to distinguish between them when writing scripts. Examine this script to see where the distinction has been made. More details regarding floating point arithmetic with computers can be found here: http://docs.python.org/tutorial/floatingpoint.html.

- Data may be normalized to total protein content by running the script called 3_normalize_to_protein.py on the internal standard-normalized data file. A second (tab delimited) file containing the total protein amounts (in mg), must be supplied as well. Below is an example:

| Peak | N2 #1 | N2 #2 | N2 #3 | eat-2 #1 | eat-2 #2 |
| --- | --- | --- | --- | --- | --- |
| Protein | 12.41 | 12.759 | 12.732 | 13.62 | 16 |

Run the script 3_normalize_to_protein.py using the example protein file sample_protein_data.txt.

Note: Our data has been obtained by collecting and analyzing excreted metabolites from C. elegans. Protein content refers to the total amount of protein present in the animals remaining in animals AFTER metabolites have been collected (see Methods for protein extraction procedure). It is necessary to normalize to total protein content prior to comparing data files since different C. elegans genetic mutants have different adult masses. If one wishes to apply this methodology to a different system (e.g. metabolites in rat plasma, etc.) where all metabolites are obtained from the same animal, depending on the time course over which the data is collected changes in total protein content may become irrelevant and this step could be omitted.

- Run the script called 4_remove_artifacts.py on the normalized data. This will remove known artifacts from the data file and will place them in a separate file. It requires a file of known artifacts. Artifacts arising from the derivatization procedure can be identified by derivatizing and analyzing a blank sample and looking for peaks present in both the blank and metabolite samples. A list of known artifacts using our derivatization procedure is included with the Supplementary Information of this article (100906_artifact_list.txt). If a different derivatization procedure is used then different artifacts may be generated. It is very important that blank samples are included from the very beginning of sample collection. On several occasions plasticizers have leached out of poor quality plasticware and have been detected in our samples.

- Run the script called 5_convert_IDs_to_names.py on the artifact-free data. This will convert the compound ID numbers to compound names. It requires a two column list of compound ID numbers vs. compound names. This is a template that can be used for any number of changes to the compound identifier. For example: names to compound IDs, old names to new names, names to a new format of names, etc. All that needs to be adjusted is the database contained in a file of a specified name. Use the file 2011_12_19_IDs_to_names.txt included with the Supplementary Information of this article.

- Prior to hierarchical clustering and further data mining, peaks corresponding to noise need to be removed. This script requires an auxiliary file specifying the category membership of each sample. Below is an example:

| Peak | N2 #1 | N2 #2 | N2 #3 | eat-2 #1 | eat-2 #2 |
| --- | --- | --- | --- | --- | --- |
| Categories | 1 | 1 | 1 | 2 | 2 |

Note 1: Like samples need to be grouped together

Note 2: If you do not want to filter by category, just set every category to 1

A default cutoff value of 10,000 is set in the script 6_low_abundance_cutoff.py. For a peak to pass the filter it must be present above the threshold value in all samples within at least one category. An alternative value can be declared as the third argument of the command line. The value of 10,000 is a general compromise which has been obtained by inspecting several chromatograms.

Run the script entitled 6_low_abundance_cutoff.py using the supplementary file categories_sample.txt.

A useful way to store data for future use following integration in MET-IDEA is to run the python script 1_peak_name_reformat.py on the MET-IDEA output file, and then copy and paste the values into a master Excel spreadsheet. When adding new data to the spreadsheet, make sure that the rows (metabolites) are correctly aligned. As you analyze new samples you may detect additional compounds and assign them new ID numbers. Previous samples may not have been analyzed for these metabolites, or they may not be present. Sorting the data by ID number (as opposed to RT) facilitates this process. We choose to enter a value of zero into our database to indicate that a particular metabolite was not searched for in a sample. When comparisons across samples are desired, simply remove any columns (samples) from the spreadsheet which are not to be used in the comparison, and save as a tab delimited .TXT file under a different name. The remaining scripts can then be run on the resulting file, and clustering performed. Storing data in this way makes it fast and easy to cluster different permutations of samples, and to compare samples which were analyzed at different times.

# 10. Hierarchical Clustering of Data Using GenePattern

Following data processing, one way to visualize groups of metabolites which are altered across samples is to use Hierarchical Clustering. Data can then be viewed in the form of heat maps which may be colored either globally (comparing metabolite [ion intensity] levels across the entire dataset) or locally (comparing levels a single metabolite [ion intensity] between samples). Since ion intensity does not equate with absolute metabolite concentration (Section 7), high intensity features may not correspond to high concentrations, therefore we find the relative coloring scheme to be more informative. Additionally, in this format, cluster analysis trees indicate samples which have similar features across categories. Hierarchical clustering and visualization of data is performed with GenePattern (http://www.broadinstitute.org/cancer/software/genepattern/). Analyses can be run on the public server at the Broad Institute, without need to download software.

- For this analysis, use the supplementary file MiMB_N2_eat-2.gct. To generate a .GCT file from processed MET-IDEA results, run the Python script 7_METIDEA_to_gct_format.py. This will put the data into the correct format (.GCT) for GenePattern. The output file will be formatted as in the example below.

- Go to the GenePattern website and select Run analyses on the Broad public server. You will be prompted for login information. If you do not have an account you will need to register and create one. Registration is free.

- Select Clustering

- Select Hierarchical clustering

- Go to Step 2: HierarchicalClustering and select Open HeirarchicalClustering

- In the Input filename field browse and select the processed data file (.GCT format) to be analyzed.

- In the column distance measure field select Pearson correlation. This is a measurement of the correlation between two variables, and is a measure of the strength of linear dependence between them.

- In the row distance measure field select Pearson correlation.

- In the clustering method field select Pairwise complete-linkage.

- Click on the Run button at the top of the screen.

- Once the job has finished, three files will appear. Select the arrow at the end of the filename of one of the three output files, and under the drop down menu which appears, select HierarchicalClusteringViewer.

- When the HierarchicalClusteringViewer display appears, click Run. A clustered heatmap will be displayed.

- By selecting View Options under the View menu the color scheme can be toggled between a relative and a global view and colors can be adjusted. The image can be saved in a variety of standard image formats by selecting Save Image under the File menu.

# 11. Summary - Overall Flow for Processing Metabolomics Data

Above we have presented an overview of several programs useful for processing GC-MS data. Many of these programs also have additional functionalities which fall beyond the scope of this tutorial. In this section, we will summarize by providing an overall work flow for using these programs to process metabolomics data.

- Prior to processing data, examine it with Xcalibur.

Things to look for:

- Overall intensity. How does it compare with previous samples?

- Poorly shaped peaks due to column overloading.

- Inaccurate masses due to instrument calibration error. Inaccurate masses will lead to errors in AMDIS and MET-IDEA.

- Presence of internal standards. Make a note of the retention times and masses of internal standards for double checking the MET-IDEA calibration.

- Use AMDIS to deconvolute the GC-MS chromatogram into component peaks and their respective spectra. This can be used to generate an IRT list, or to identify new components which are not in your library. An existing IRT list may be manually updated with unknown components discovered in AMDIS.

- Use MET-IDEA to integrate peaks based on your IRT list.

- Convert data files to .CDF format using the Xcalibur file convertor, or a file converter specific to your instrument, for import into MET-IDEA.

- Place of copy of the IRT list (.ION file) into the same folder as your data. You can use the file 2011_10_24_IRT.ION included with the supplementary files. This file is Kovats Retention Time indexed, and it is recommended to keep this as a master file. When performing analyses, it is sometimes helpful to adjust the retention times in the IRT to correct for instrument drift. This is easily done in Microsoft Excel by adding or subtracting a constant value to all RTs so that the RT of the internal standard is correct. This should be saved as a working file for the current dataset which you are processing.

- Start MET-IDEA and double check the parameters.

Useful parameters are listed below but these may be modified to suit your data:

[Chromatography = GC]

AVG_PEAK_WIDTH = 0.10

MIN_PEAK_WIDTH = 0.5

MAX_PEAK_WIDTH = 1.1

PEAK_START_STOP_SLOPE = 2.5

ADJUST_RT_ACCURACY = 0.25

PEAK_OVERLOAD_FACTOR = 0.75

[MASS_SPEC = Quadrupole]

MASS_ACCURACY = 0.1

MASS_RANGE = 0.4

[AMDIS]

EXCLUDE = 73 147 281 341 415

MASS_LIMIT = 150

- Start the MET-IDEA processing and navigate to the .ION file in the data folder.

- Select calibration settings.

- Select calibration peaks. These can be the internal standards, or any peak(s) which are present in all of your samples. For our samples we use the internal standards norvaline (ID 100145, m/z 186.17, RT 26.32 min), phenylpyruvate (ID 100144, m/z 250.1, RT 29.21 min), and 3,4-dimethoxybenzoic acid (ID 100080, m/z 239.02, RT 32.54 min).

- Evaluate the calibration results (.CAL file). Do they agree with the values you recorded when examining the data with Xcalibur? If the values are too far off, MET-IDEA won't find the reference peaks and values will need to be manually added for the calibration. Enter the value that would need to be added or subtracted from the data file to match the ion-retention time list. Save the calibration file. Note: if the MET-IDEA results show an absence of peaks then most likely the calibration is wrong.

- Save the integration results and reopen. Examine the results and look for integration errors. Correct major errors by manually entering peak start/stop times. If there are a lot of errors, you may need to reevaluate your calibration or parameter settings.

- Transpose the file and save changes. If you have access to the raw data files on your instrument, and the software to convert them to .CDF format, you may wish to delete the .CDF files at this point. The .CDF files are approximately 50% larger than Xcalibur .RAW files, and may be as large as 15-18 MB each.

- Process the MET-IDEA results using Python

- Make sure that the result file is in the proper format: Peak in rows, samples in columns.

- Copy the processing scripts, artifact list and IDs_to_names file into the same folder as your data. Open a command window in you data folder.

- Run 1_peak_name_reformat.py (Open scripts in a text editor for details on input requirements).

- Run 2_normalize_to_IS.py (You can specify the internal standard or leave as the default: 3,4-dimethoxybenzoic acid).

- Run 3_normalize_to_protein.py (You will need a file with the total protein amounts).

- Run 4_remove_artifacts.py (You will need a file with a list of known artifacts).

- Run 5_convert_IDs_to_names.py (You will need a file with the IDs and associated compound names).

- (Optional) Data can be filtered to remove noise peaks below a certain threshold by running 6_low_abundance_cutoff.py with an arbitrary cutoff which you determine to be appropriate.

- Submit data to GenePattern for hierarchical clustering and visualization

- Run 7_METIDEA_to_gct_format.py to convert processed MET-IDEA data to .GCT format

- Submit to GenePattern  HierarchicalClustering.

- Use the HeirchicalClusteringViewer to visualize the clustered data and generate a heat map.

PAGE   \*